In vitro Assessment of Tissue Damage Due to Neural Prosthetic Device Insertion

CS Bjornsson, SJ Oh, Y Al-Kohafi, LJ Lim, KL Smith, JN Turner, S De, B Roysam, W Shain, SJ Kim

Wadsworth Center, New York State Dept. of Health, Albany, NY

Seoul National University, Seoul, Korea

Rensselaer Polytechnic Institute, Troy, NY

The potential benefits of neuroprosthetic devices are currently challenged by biological reactive responses that interfere with the brain-machine interface. Reactive responses begin immediately upon insertion; however the contributions of many of the physical parameters of insertion, which may determine how much damage is caused, are presently unknown. We have developed an in vitro system that allows us to capture time-lapse images of tissue deformation during neuroprosthetic device insertion, by labeling the vasculature of live coronal slices, and used this model to compare the damage caused by inserting devices with several tip shapes at different insertion speeds. Brain slices labeled to highlight neurovasculature permitted observation of specific vascular damage events, including displacement of luminal contents, breakage, rupture, and vessel dragging, that may cause collateral damage to the surrounding nervous tissue. Damage to transcerebral arteries at times occurred over 100 microns from the insertion site, suggesting that the extent of the reactive responses may depend on initial damage to the neurovasculature as well as intercellular signaling and diffusion of soluble factors. Using automated image analysis we could track one hundred interest points within the tissue to measure tissue displacement and maximum and effective strains. Slower insertion speeds resulted in greater displacement and strain. Device tip shape was less influential, and tips of all shapes caused comparable tissue displacement and strain as well as breakage, rupture and dragging of neurovasculature. Surprisingly, cortical surface features were one of the most influential factors affecting insertion damage; attempts to insert devices through pial blood vessels resulted in severe compression and damage. In addition, pial vessels were often carried into the neural parenchyma with the device. The results of this study will help guide future design of neuroprosthetic devices.